## Slide 1: ISO/IEEE 11073 MDDLHL7 FHIR Transformational Methodology and Applications - Overview
HL7 Pt. Care DCM4MD, GAS, FHIR and IEEE 11073 Working Groups
12Jan2013
FHIR-MDDL_Schema_jw7a.pptx
This document provides an overview of ISO/IEEE 11073 MDDL to HL7 FHIR transformational methodology and applications, such as device-to-”enterprise” interoperability in general and non-invasive blood pressure (NIBP) data in CDA format as part of an Anesthesia Record in particular, in support of HL7 Project 850 concerning CDA transform of NIBP, which provides the primary use case (UC).
Content
  Figure 1. Topological model
  1A. Transformation model
  1B. Distribution model
  Figure 2. Development model
  Appendices
  A Use Case (UC) modeling
  A1. Natural language to synthesized relational transform
  A2. Synthesized MDDL to FHIR relational transform
  B Device profile Library schema
  C Model of models (MoM)
  D Entropy / System of [sub-]systems (SoS)model
  E MDDL FHIR Resource Analogues

## Slide 2
MED DEV (Mgr)
UML META-MODEL
RECORD
.VITAL_SIGNS . PERI-OP.ANES
[CDA]
(Content)
MDDL DIM-based model
[XSLT/]XML CDA TRANSFORM
IEEE 11073 103xx/10201 DIM Standard Model-based (Content]
FHIR-based XML TRANSFORM
e.g. @ PHMR template
 | FHIR | Resource |  | DEVICE | .MDS | .Simple | 
 | FHIR | Resource |  | Metric | .NUmeric |  | 
@ MDDL  RMIM transform/map
FHIR RESOURCE
(MDDL METRIC.NU
FHIR RESOURCE
(MDDL MDS)
Mgr MDIB
MDAP Mini-ASS’N
e.g. JSON
MDAP Primary ASS’N
Agent MDIB
MDAP Primary ASS’N
Proxy MDIB
TPORT Profile
TPORT Profile
MDDL2CDA FHIR-based TRANSFORMER
FHIR Proxy
e.g. JSON
TPORT Profile
CDA RESOURCE
TPORT Profile
Indirect config (i.e. Gatewayed)
Direct config (i.e. Dev2EHRIS)
B) Distribution model. “Lower” model
A) Transformation model. “Upper” model
ISO OSI Reference
Appl’n layer
Present’n/Session layers
Transport-//PHY layers
Proxy in Agent vs. Mgr
May be further distributed @ CIS, HIS network architecture
Figure 1. Topological model
Standalone config
Transformer in Agent vs. Mgr or networked CIS or HIS
MED DEV (Agent)
DIM MODEL
CHANnel
MedDevSys (MDS)
May be composite
VirtualMedDev (VMD)
PHD does not use VMD or CHAN
Metric
(.NUmeric)
Other Metrics:
.ENUMeration
.NU.Cmpd
.NU.Cmplx
.RTSA
.CSA
≥1
≥1
≥1
MDDL CLASSes
See SysOfSys model

## Slide 3
UML META-MODEL
DIM MODEL
MDS
*
[VMD]
*
[CHAN]
*
METRICs
MDDL DIM-based model
FHIR-based XML TRANSFORM
 | FHIR | Resource |  | DEVICE | .MDS | .Simple | 
 | FHIR | Resource |  | Metric | .NUmeric |  | 
@ MDDL  RMIM transform/map
FHIR RESOURCE
(MDDL METRIC.NU
FHIR RESOURCE
(MDDL MDS)
Figure 2. Development model
Synthesize MDS relational model from Use Case (UC) and enter into DIM Editor
Export
(and adapt as needed) into
x73_FHIR-DeviceProfile .SimpleMDS .NIBP @ GASx73
GASx73 NIBP USE CASE
x73-_FHIR-DevProfile Library
Synthesized Use Case relational model
CDA

## Slide 4
Appendix A. Use Case (UC) modeling
A1. Natural language to synthesized relational transform
Initially, the Use Case is coded into a relational analogue of x73 MDDL. This case models a 3-parameter NIBP device as a “Simple MDS”, which has a single VMD, and a single Channel. See Appendix A2 for relational spreadsheet form.
Use Case
Patient is a 78 year old woman undergoing a total left hip replacement is placed in a right lateral decubitus position with her left arm out on a board that has an angle of elevation of 0° and is at an angle of 90° to the operating table.
Probes and measurement devices are connected to a patient so that her physiological values can be measured and recorded during an intra-operative procedure. Relevant to this example, a blood pressure cuff is placed on the upper left arm.
The cuff is attached to an automatic non-invasive blood pressure measuring module (Patmon BP 900, S/N: 45678, Firmware Rev 1.2.1) which is fitted in a Patmon 8760 multi-function patient monitor ( S/N 12345, Firmware Rev 1.2.3).
The Patmon 800 BP module uses a sphygmomanometric technique to measure systolic and diastolic blood pressure. Mean arterial pressure (MAP), calculated by the monitor, is approximated using the following formula:
MAP = DP + 1/3(SP – DP)
… where MAP is Mean Arterial Pressure, DP is diastolic pressure and SP is Systolic Pressure
All blood pressure values are expressed in millimeters of Mercury (mmHg).
Monitoring is started at 10:00:00 and the set of measurements (DP, SP and calculated MAP) are available at 10:02:00. Thereafter, the same set of measurements is available at 5 minute intervals. Monitoring is discontinued at 12:30:00.
During the procedure the cuff is loosened and re-fixed in the same position at 10:37:00. This necessitates the loosening of bands that secure the patient’s arm to the board and their subsequent replacement.
Patient
Demographics
Age=78 year old
Gender=woman
undergoing a
ClinicalProcedure
Type total left hip replacement
Positioning.Interface is placed in a right lateral decubitus position with her left arm out on a board that has an angle of elevation of 0° and is at an angle of 90° to the operating table.
MDS [/Probes and measurement devices]
are connected to a patient so that her physiological values can be measured and recorded during an intra-operative procedure. Relevant to this example, a
VMD
Type [non-invasive] blood pressure
[Cuff]Site is placed on the upper left arm. The
cuff is attached to an
Channel automatic non-invasive blood pressure measuring module
Type (Patmon BP 900, S/N: 45678,
Firmware Rev 1.2.1) which is fitted in a
MDS
Model Patmon 8760 multi-function patient monitor (
S/N 12345,
Firmware Rev 1.2.3).
Channel The Patmon 800 BP module
MeasurementMethod uses a sphygmomanometric technique to measure
<continued in next column>
Metrics
Class.Type: Numeric
. systolic blood pressure.
.diastolic blood pressure
.Mean arterial pressure (MAP),
DerivationComponent calculated by the monitor,
DerivationFormulation is approximated using the following formula:
MAP = DP + 1/3(SP – DP)
… where MAP is Mean Arterial Pressure, DP is diastolic pressure and SP is Systolic Pressure
UnitsOfMsmt All blood pressure values are expressed in millimeters of Mercury (mmHg).
DataAcquisition
Start time Monitoring is started at 10:00:00 and the set of measurements (DP, SP and calculated MAP) are available at 10:02:00.
SamplingPeriodicity Thereafter, the same set of measurements is available at 5 minute intervals.
End time Monitoring is discontinued at 12:30:00.
ClinicalProcedure
[Re-]Positioning.Interface During the procedure the cuff is loosened and re-fixed in the same position at 10:37:00. This necessitates the loosening of bands that secure the patient’s arm to the board and their subsequent replacement.

## Slide 5
 | FHIR | Resource |  | DEVICE | .MDS | .Simple | 
 | FHIR | Resource |  | Metric | .NUmeric |  | 
@ MDDL  RMIM transform/map
FHIR RESOURCE
(MDDL METRIC.NU
FHIR RESOURCE
(MDDL MDS)
Appendix A2. Synthesized MDDL to FHIR relational transform
The initial relational coding (see Appendix A1) is further refined using a spreadsheet tool, yielding the model at left. The general correspondence to MDDL MDS FHIR relational form is shown at right. FHIR essentially layers the HL7 RMIM previously developed for the GASx73 CDA project, which is shown in the background.

## Slide 6
Appendix B. Device Profile Library schema
FHIR “Device profile” .xml was imported into an MS Access DB and resulting tables (at left), edited, and a Query developed to yield the result (at bottom); each row instantiates a “FHIR Device profile” according to the general profile attributes in the FHIR import and the particular attributes entered into the “Definitions Profile-cases” table, yielding the result (at bottom).

## Slide 7
UML META-MODEL
http://dim.prometheuscomputing.com/gb_users/login
CHANnel
MedDevSys (MDS)
May be composite
VirtualMedDev (VMD)
PHD does not use VMD or CHAN
Metric
(NUmeric)
Other Metrics:
.ENUMeration
.NU.Cmpd
.NU.Cmplx
.RTSA
.CSA
MDS
TYPEs
≥1
≥1
≥1
MDDL CLASSes
RECORD
.VITAL_SIGNS . PERI-OP.ANES
[CDA]
(Content)
MDDL DIM
For DIM Editor “Demo” :model, see
CONFIG . DEMO
FHIR RESOURCES map salient x73 constructs, layering HL7 v3 RIM constructs
XML-->[XSLT] TRANSFORM
IEEE 10201 DIM Standard (Content]
XML TRANSFORM
 | FHIR | Resource |  | DEVICE | .MDS | .Simple | 
 | FHIR | Resource |  | Metric | .NUmeric |  | 
as substitution for direct X73-RMIM-CDA xform
e.g. @ PHMR template
Assumes that FHIR layers RIM; RMIM may be mapped @ FHIR layer
Appendix C. Model of Models - Overview
DCM Use Case (UC): see MH20120102 – Non invasive BP use Case.docx
GASx73 RMIM CDA overview: see GASx73-CDA-DCM-DataTopo-UCs-DCM_d1-jw

## Slide 8
RT Display
Entropy considers the dynamics of information integration. The “upper” model defines the ‘types’ while the ‘lower model considers distribution of “workgroup”-oriented [sub-]system clusters that may be relatively flexibly configured dynamically.
Appendix D. Entropy/System of [sub-]systems model
RECORD
.[CDA]
(Content)
.ANES REC
.FLOWSHT
ACM Event
ACM Event
.PHR
Acuity
[Peri-Op]
Critical
Acute
Sub Acute
Emerg /Evac
Chronic
Home | Evac | In-Patient |  |  | e-Patient
 |  | Sub-Acute | Acute | Critical | [CDS/CIS
PHD-AHD | DefibMon /{*PHD} | {PWD{PoC,Cent Mon}} |  | {{[Anes]MonWkstn} |  | Wkgrp
Per Pt./Intermittent | Per Pt./Session | Multi-patient |  |  | Per Pt./Procedure | Multi-Pt./Consult
Weight; BP, Meds Glucose, Pt. Locn | Vitals/Cts, EKG | Vitals, other Cts RT/Spot Near RT | Vitals, other/Cts, NearRT Info |  |  | Controls
PHD[wB/PAN] |  | wB/P/LAN | wLAN | wB/P/LAN |  | wL[/W]AN
Continua [w]WAN |  | Telemetry/SRR RealtIme Cts [PCD-0n; JSON ?] |  |  |  | [PCD-??; DDS ?]
IMAGERY
PHD-104xx
PoC-103xx
PoC-103xx GAS x73 CDA
Central Mon
Shared Pt:Dev Bindings (Pairing)
PoC-103xx
PoC MonM
TelePortable [Defib] Mon
Pt.WornDevMon
Anes/Vent/Mon
PCA
CDS Mon
Home Mon
Pt. Trf.
PCA
ORDER
Return to Topo
RT Display
Pt.
RT ALG
RT ALG
RT Display
Event “Data” “RT” “Doc”

## Slide 9
Appendix E. MDDL FHIR Resource Analogues
x73-_FHIR-DevProfile Library  DIRectory of Device Profile Instances
MDDL “Event” info structures are a good candidate for mapping transforms to FHIR. In the following diagram, an “EVENT LOG” comprises a [time-stamped] list of Events Info constructs, typically NUMerical OBSERVATIONS; however, there are a variety of EventLogEntry types and OBS types including ENUMerations, ALERTS, and even CONTROL/operation reports.
Extensions
-EVENT LOG
…
-CONTROL PKG
EVENT LOG
SCAN REPORT INFO
[NUMeric] OBSERVATION